Methods for the Collection and Measurement of Chromophoric Dissolved Organic Matter (CDOM) and Dissolved Organic Carbon (DOC)

Reference for DOC and CDOM methods:

Mannino, A., M. E. Russ, and S. B. Hooker (2008), Algorithm development and validation for satellite-derived distributions of DOC and CDOM in the U.S. Middle Atlantic Bight, J. Geophys. Res., 113, C07051, doi:10.1029/2007JC004493.

Samples for determination of CDOM spectral absorption coefficients are stored under refrigeration.  In the laboratory, CDOM samples are warmed to room temperature and filtered through 0.2 µm Whatman Nuclepore polycarbonate filters or Gelman Supor (polyethersulfone) filters prior to analysis (Mitchell et al. 2000).  Filtration of CDOM samples through GF/F filters is necessary for coastal ocean waters due to the high particle load that quickly clog Nuclepore and Supor filters.  Absorption spectra of CDOM are measured using a Cary 100 UV-Visible scanning spectrophotometer (250-800 nm) with UV oxidized Milli-Q water as the blank and reference (Mitchell et al. 2003). Instrument scan settings were as follows:  250–800 nm wavelength scan range, 1 nm data interval, 100 nm min-1 scan rate, and 4 nm slit width. The instrument noise for reference baselines of air-to-air and ultrapure water spectral scans was within ±0.0005 absorbance units. Spectral absorption coefficients are determined after subtracting the raw absorption measurements with field filtration blanks of UV-oxidized Milli-Q and a null point value an average of absorbance for 695-700nm (modified from Mitchell et al. 2000; 2003).  Where absorption spectra of filtration blanks are within the level of instrument noise, additional corrections are not necessary.  Instrument performance tests (wavelength accuracy and reproducibility, photometric noise, and baseline flatness) were conducted each day prior to analysis. Furthermore, National Institute of Standards and Technology (NIST)-traceable calibration standards (Holmium oxide filter for wavelength accuracy and Spectronics standards, Thermo Electron Corporation, to evaluate stray light, wavelength accuracy, and photometric performance) were also used to verify instrument performance. The uncertainty associated with CDOM spectral absorption coefficients at an instrument noise level <0.0046 m-1 was on the order of 0.023 to 0.039 m-1 and based on the summation of the instrument manufacturer’s guaranteed specifications for photometric accuracy, stability and noise. Repeat scans of seawater subsamples, replicate subsamples, and Niskin versus persistaltic pump samples yielded a coefficient of variation of typically <±3%. The CDOM spectral slope is determined by fitting a single-exponential non-linear curve to the measured curve from 300-600nm [a(λ)=a(λo)e-S(λ- λo), where a(λ) and a(λo) represent the absorption coefficients at wavelength λ and reference wavelength λo].   CDOM absorption coefficients, aCDOM(λ), and the spectral slope coefficient (S), a parameter that quantifies the exponential absorption decrease with increasing wavelength, have been shown to vary with type and source of CDOM.  The CDOM spectral slopes derived from this non-linear curve fit are insensitive to aCDOM from higher wavelengths; i.e., equation affords greater weight to aCDOM values in the ultraviolet and blue spectral regions where signal-to-noise is highest.

Benner, R. and M. Strom. 1993. A critical evaluation of the analytical blank associated with DOC measurements by high-temperature catalytic oxidation. Mar. Chem. 41: 153-60.

Blough, N.V. and R. Del Vecchio. 2002. Chromophoric DOM in the coastal environment, pp.

509-546. In:  Hansell D.A. and Carlson C.A. (eds), Biogeochemistry of Marine Dissolved Organic Matter.  Academic Press.

Mitchell, B.G., et al. 2000. Determination of spectral absorption coefficients of particles, dissolved material and phytoplankton for discrete water samples, pp. 125-153. In: Fargion G.S. and Mueller J.L. (eds), Ocean Optics Protocols for Satellite Ocean Color Sensor Validation. NASA/TM-2000-209966.

Mitchell, B.G., M. Kahru, J. Wieland and M. Stramska. 2003. Determination of spectral absorption coefficients of particles, dissolved material and phytoplankton for discrete water samples, pp. 39-64. In: Mueller J.L.,  G.S. Fargion and C.R. McClain (eds), Ocean Optics Protocols for Satellite Ocean Color Sensor Validation. NASA/TM-2003-211621/Rev4-Vol.IV.

Sharp, J.H., C.A. Carlson, E.T. Peltzer, D.M. Castle-Ward, K.B. Savidge and K.R. Rinker. 2002. Final

dissolved organic carbon broad community intercalibration and preliminary use of DOC reference materials. Mar. Chem. 77: 239-253.